## Settings
Settings für die Validierung
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Institution / Labor | CVUA S
Parameter | Gattung (Genus) |  | Art (Spezies)
|  | Vaccinium | myrtillus
Unterscheidungsgrenzen
| Score zur Unterscheidung der Gattung | 1.7 | Score zur Unterscheidung der Spezies | 2

## Parameter (Spezies)
Parameter: | Vaccinium | myrtillus | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | nicht zugeordnet | richtig-positiv | falsch-negativ | falsch zugeordnet
|  |  | 1 | 1 | TX 637 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | myrtillus | 2.09 | Vaccinium | myrtillus | 2.09 | A | 1 | 0 | 1 | 0 | 0
Zahl aller Spektren | 22 |  | 1 | 1 | TX 192.3 | 2; expert opinion | Plant extr.-R | 2022-12-06 | Vaccinium | myrtillus | 2.39 | Vaccinium | myrtillus | 2.15 | A | 1 | 0 | 1 | 0 | 0
davon Identifiziert | 22 |  | 1 | 1 | TX 469.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-22 | Vaccinium | myrtillus | 2.08 | Vaccinium | myrtillus | 2.07 | A | 1 | 0 | 1 | 0 | 0
nicht identifiziert: | 0 |  | 1 | 1 | TX 507.2 | 2; label; sensory analysis | Plant extr.-RL40 | 2022-12-02 | Vaccinium | myrtillus | 2.17 | Vaccinium | myrtillus | 2.17 | A | 1 | 0 | 1 | 0 | 0
von allen identifizierten Isolaten |  |  | 1 | 1 | TX 550.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | myrtillus | 2.32 | Vaccinium | myrtillus | 2.32 | A | 1 | 0 | 1 | 0 | 0
richtig-positiv: | 22 |  | 1 | 1 | TX 571.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-12 | Vaccinium | myrtillus | 2.1 | Vaccinium | myrtillus | 2.09 | A | 1 | 0 | 1 | 0 | 0
falsch zugeordnet | 0 |  | 1 | 1 | TX 572.2 | 2; label; sensory analysis; Biomarker | Plant extr.-R | 2022-12-23 | Vaccinium | myrtillus | 2.03 | Vaccinium | myrtillus | 1.94 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 607.2 | 2; label; sensory analysis | Plant extr.-R | 2023-01-19 | Vaccinium | myrtillus | 2.07 | Vaccinium | myrtillus | 2.04 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 622 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.35 | Vaccinium | myrtillus | 2.17 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 623 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.11 | Vaccinium | myrtillus | 2.1 | A | 1 | 0 | 1 | 0 | 0
Scorewertgrenzen: | 2 | Species | 1 | 1 | TX 624 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.66 | Vaccinium | myrtillus | 2.13 | A | 1 | 0 | 1 | 0 | 0
| 1.7 | Genus | 1 | 1 | TX 625 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.23 | Vaccinium | myrtillus | 2.14 | A | 1 | 0 | 1 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 1 | 1 | TX 626 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.39 | Vaccinium | myrtillus | 2.29 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 627 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.33 | Vaccinium | myrtillus | 2.26 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 628 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.32 | Vaccinium | myrtillus | 2.28 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 629 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.44 | Vaccinium | myrtillus | 2.31 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 630 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.23 | Vaccinium | myrtillus | 2.21 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 631 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.38 | Vaccinium | myrtillus | 2.35 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 632 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.14 | Vaccinium | myrtillus | 2.13 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 633 | 3; botanical collection | Plant extr.-R | 2023-01-20 | Vaccinium | myrtillus | 2.45 | Vaccinium | myrtillus | 2.23 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 635 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | myrtillus | 2.43 | Vaccinium | myrtillus | 2.41 | A | 1 | 0 | 1 | 0 | 0
|  |  | 1 | 1 | TX 636 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | myrtillus | 2.31 | Vaccinium | myrtillus | 2.29 | A | 1 | 0 | 1 | 0 | 0
davon mit ID-Ergebnis | 22
Positive Übereinstimmungen (PA) | 22
Negative Abweichungen (ND) | 0
nicht identifiziert in % | 0%
falsch zugeordnert in % | 0%

## #Parameter (Spezies) DB
Parameter: | Vaccinium myrtillus |  | Eigenspektrum | Eigen ID | Codierung | Validität | Methode | Datum | verwendete Gattung | verwendete Art | Gattung | Art | Score | Gattung | Art | Score | Konsistenz | Identifiziert | richtige ID | keine richtige ID | 1. Hit | 2. Hit | 1. Hit Treffer Parameter | 2. Hit Treffer Parameter | falsch als Parameter identifiziert | PM unter den ersten 2 Hits
|  |  | 1 | 1 | TX 600.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes | rubrum | Ribes | rubrum | 2.12 | Ribes | rubrum | 2.04 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
Anzahl Isolate: | 41 |  | 1 | 1 | TX 003 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus | precatorius | Abrus | precatorius | 2.08 | Abrus | precatorius | 1.78 | A | 1 | 1 | 0 | Abrus | Abrus | 0 | 0 | 0 | 0
davon Identifiziert | 39 |  | 1 | 1 | TX 006 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus | precatorius | Abrus | precatorius | 2.06 | Abrus | precatorius | 1.93 | A | 1 | 1 | 0 | Abrus | Abrus | 0 | 0 | 0 | 0
von allen identifizierten Isolaten |  |  | 1 | 0 | TX 443.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | barbarum | Lycium | barbarum | 2.34 | Lycium | europaeum | 2.11 | B | 0 | 0 | 1 | Lycium | Lycium | 0 | 0 | 0 | 0
richtige ID | 39 |  | 1 | 1 | TX 447.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | chinense | Lycium | chinense | 2.48 | Lycium | europaeum | 1.17 | A | 1 | 1 | 0 | Lycium | Lycium | 0 | 0 | 0 | 0
falsch als Parameter identifiziert | 0 |  | 1 | 1 | TX 449.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium | europaeum | Lycium | europaeum | 2.13 | Lycium | chinense | 1.41 | A | 1 | 1 | 0 | Lycium | Lycium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 599.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes | nigrum | Ribes | nigrum | 2.33 | Ribes | nigrum | 2.06 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 603.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-21 | Ribes | nigrum | Ribes | nigrum | 2.5 | Ribes | nigrum | 2.21 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 584.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.36 | Ribes | rubrum | 2.26 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 585.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.42 | Ribes | rubrum | 2.41 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
Parameter wird getroffen mit: |  |  | 1 | 1 | TX 587.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.4 | Ribes | rubrum | 2.36 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 588.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.48 | Ribes | rubrum | 2.4 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
1. Hit Treffer Parameter | 0 |  | 1 | 1 | TX 589.2 | 2; label; sensory analysis | eDT | 2023-02-08 | Ribes | rubrum | Ribes | rubrum | 2.33 | Ribes | rubrum | 2.23 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
2. Hit Treffer Parameter | 0 |  | 1 | 1 | TX 590.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.36 | Ribes | rubrum | 2.31 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
PM unter den ersten 2 Hits | 0 |  | 1 | 1 | TX 591.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.28 | Ribes | rubrum | 2.11 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
in % | 0% |  | 1 | 1 | TX 593.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.46 | Ribes | rubrum | 2.39 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 594.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes | rubrum | Ribes | rubrum | 2.23 | Ribes | rubrum | 1.96 | A | 1 | 1 | 0 | Ribes | Ribes | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 575.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium | macrocarpon | Vaccinium | macrocarpon | 2.35 | Vaccinium | myrtillus | 1.13 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
Scorewertgrenzen: | 2 | Species | 1 | 0 | TX 576.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium | macrocarpon | Vaccinium | macrocarpon | 1.98 | Vaccinium | sect. | 0.94 | B | 0 | 0 | 1 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
| 1.7 | Genus | 1 | 1 | TX 468.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-22 | Vaccinium | sect. | Vaccinium | sect. | 2.34 | Vaccinium | sect. | 2.2 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
Es wird die gesamte zu überprüfende Datenbank verwendet. |  |  | 1 | 1 | TX 470.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-12 | Vaccinium | sect. | Vaccinium | sect. | 2.38 | Vaccinium | sect. | 2.32 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 492.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-21 | Vaccinium | sect. | Vaccinium | sect. | 2.41 | Vaccinium | sect. | 2.3 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
davon mit ID-Ergebnis | 39 |  | 1 | 1 | TX 501.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.22 | Vaccinium | sect. | 2.01 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
Negative Übereinstimmungen (NA) | 39 |  | 1 | 1 | TX 502.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.42 | Vaccinium | sect. | 2.38 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
Positive Abweichungen (PD) | 0 |  | 1 | 1 | TX 503.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.5 | Vaccinium | sect. | 2.43 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 504.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.36 | Vaccinium | sect. | 2.18 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
nicht identifiziert in % | 4.878% |  | 1 | 1 | TX 505.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium | sect. | Vaccinium | sect. | 2.27 | Vaccinium | sect. | 2.25 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
falsch zugeordnert in % | 0% |  | 1 | 1 | TX 506.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.46 | Vaccinium | sect. | 2.17 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 508.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.43 | Vaccinium | sect. | 2.34 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 509.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.4 | Vaccinium | sect. | 2.22 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 510.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium | sect. | Vaccinium | sect. | 2.32 | Vaccinium | sect. | 2.31 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 511.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.31 | Vaccinium | sect. | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 512.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.4 | Vaccinium | sect. | 2.34 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 513.2 | 2; label; sensory analysis | Plant extr.-RL40 | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.3 | Vaccinium | sect. | 2.26 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 514.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium | sect. | Vaccinium | sect. | 2.28 | Vaccinium | sect. | 2.24 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 573.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium | sect. | Vaccinium | sect. | 2.21 | Vaccinium | sect. | 1.92 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 604.2 | 1; label | Plant extr.-R | 2023-01-19 | Vaccinium | sect. | Vaccinium | sect. | 2.26 | Vaccinium | sect. | 1.82 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 638 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | uliginosum | Vaccinium | uliginosum | 2.02 | Vaccinium | myrtillus | 1.29 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 639 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.15 | Vaccinium | vitis-idaea | 2.14 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 640 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.4 | Vaccinium | vitis-idaea | 2.13 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0
|  |  | 1 | 1 | TX 641 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium | vitis-idaea | Vaccinium | vitis-idaea | 2.29 | Vaccinium | vitis-idaea | 2.05 | A | 1 | 1 | 0 | Vaccinium | Vaccinium | 0 | 0 | 0 | 0

## #Parameter (Spezies)
Parameter: | Vaccinium | myrtillus | Codierung | Validität | Methode | Datum | verwendeter Organismus | Score | richtig-negativ | falsch-positiv
Test auf spektrale Nähe zum Parameter |  |  | TX 600.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes rubrum | 0.48 | 1 | 0
Anzahl Isolate: | 41 |  | TX 003 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus precatorius | 0.5 | 1 | 0
|  |  | TX 006 | 3; determination by expert | Plant extr. | 2021-07-22 | Abrus precatorius | 0.39 | 1 | 0
richtig-negativ: | 41 |  | TX 443.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium barbarum | 0.78 | 1 | 0
|  |  | TX 447.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium chinense | 0.25 | 1 | 0
falsch-positiv: | 0 |  | TX 449.3 | 3; botanical collection | Plant extr.-R | 2022-11-15 | Lycium europaeum | 0.78 | 1 | 0
|  |  | TX 599.2 | 1; label | Plant extr.-R | 2022-12-21 | Ribes nigrum | 0.28 | 1 | 0
| 0% |  | TX 603.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-21 | Ribes nigrum | 0.38 | 1 | 0
|  |  | TX 584.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.24 | 1 | 0
|  |  | TX 585.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.32 | 1 | 0
Scorewertgrenzen: | 2 | Species | TX 587.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.12 | 1 | 0
| 1.7 | Genus | TX 588.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.09 | 1 | 0
Es wird gegen eine nur aus dem Parameter bestehende DB geprüft. |  |  | TX 589.2 | 2; label; sensory analysis | eDT | 2023-02-08 | Ribes rubrum | 0.05 | 1 | 0
|  |  | TX 590.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.59 | 1 | 0
|  |  | TX 591.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.21 | 1 | 0
|  |  | TX 593.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.2 | 1 | 0
|  |  | TX 594.2 | 2; label; sensory analysis | Plant extr.-R | 2023-02-06 | Ribes rubrum | 0.55 | 1 | 0
|  |  | TX 575.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium macrocarpon | 1.13 | 1 | 0
|  |  | TX 576.2 | 1; label | Plant extr.-R | 2022-12-19 | Vaccinium macrocarpon | 0.88 | 1 | 0
|  |  | TX 468.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-22 | Vaccinium sect. | 1.1 | 1 | 0
|  |  | TX 470.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-12 | Vaccinium sect. | 1.06 | 1 | 0
|  |  | TX 492.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-21 | Vaccinium sect. | 1.16 | 1 | 0
|  |  | TX 501.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 1.4 | 1 | 0
|  |  | TX 502.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 1.21 | 1 | 0
|  |  | TX 503.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 1.27 | 1 | 0
|  |  | TX 504.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 1.17 | 1 | 0
|  |  | TX 505.2 | 2; label; sensory analysis | Plant extr.-AceR | 2022-11-29 | Vaccinium sect. | 1.32 | 1 | 0
|  |  | TX 506.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 1.12 | 1 | 0
|  |  | TX 508.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 1.21 | 1 | 0
|  |  | TX 509.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 1.02 | 1 | 0
|  |  | TX 510.2 | 2; label; sensory analysis | Plant extr.-R | 2022-11-30 | Vaccinium sect. | 1.23 | 1 | 0
|  |  | TX 511.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium sect. | 1.36 | 1 | 0
|  |  | TX 512.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium sect. | 1.11 | 1 | 0
|  |  | TX 513.2 | 2; label; sensory analysis | Plant extr.-RL40 | 2022-12-02 | Vaccinium sect. | 1.15 | 1 | 0
|  |  | TX 514.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-02 | Vaccinium sect. | 1.26 | 1 | 0
|  |  | TX 573.2 | 2; label; sensory analysis | Plant extr.-R | 2022-12-16 | Vaccinium sect. | 1.22 | 1 | 0
|  |  | TX 604.2 | 1; label | Plant extr.-R | 2023-01-19 | Vaccinium sect. | 1.13 | 1 | 0
|  |  | TX 638 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium uliginosum | 1.29 | 1 | 0
|  |  | TX 639 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium vitis-idaea | 1.32 | 1 | 0
|  |  | TX 640 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium vitis-idaea | 1.56 | 1 | 0
|  |  | TX 641 | 3; botanical collection | Plant extr.-R | 2023-01-30 | Vaccinium vitis-idaea | 1.3 | 1 | 0

## Report (Spezies)
Zielgerichtete Identifizierung
nach den Leitlinien zur MALDI-Validierung der §64 AG MALDI am BVL, 12.04.2021
Parameter: Vaccinium myrtillus
Version: | 1
erstellt von: | Dyk | geprüft von: | Rau
erstellt am: | 2023-02-07 | geprüft am: | 2023-02-14
Anwendung | Pflanzenmaterialien (Samen), auch aus Lebensmitteln
Umfeld | unverarbeitete und verarbetete Lebensmittel mit Fruchtanteilen
Verwendete Datenbank / Programmversion |  | Pflanzen DB (196 MSPs) | BT 4.0
Validierungsset: | andere Pflanzenmaterialien incl. Früchte und Beeren
Validierungsisolate/-materialien (Parameter): 22 |  | Vergleichsisolate/-materialien (#Parameter): 41
Kultivierungsbedingungen bei Mikroorganismen: -
Zustand der Probe bei sonstigen Materialien: gefroren, frisch, getrocknet
Aufreinigungsmethode: Plant extr.-R
Es werden nur identifizierte Proben für die Ermittlung der In- und Exklusivität herangezogen.
Richtig-positiv: 1. Hit =Ziel-Parameter mit Score ≥ 2 und Score 2. Hit < 1.7 oder Score 2. Hit ≥ 1.7 & Gattung gleich 1. Hit oder Score 2. Hit ≥ 2 & Art gleich 1. Hit |  | Richtig-negativ: 1. Hit ohne Zuordnung zum Ziel-Parameter mit score ≥ 2
Falsch-negativ: Kriterien für richtig-positiv werden nicht erfüllt |  | Falsch-positiv: Kriterien für richtig-negativ werden nicht erfüllt
Validierungsergebnis | Ziel-Parameter: | Vaccinium myrtillus |  | SOLL | erfüllt
Zahl aller Spektren | 22
davon mit ID-Ergebnis | 22 | Identifiziert Rate | 100% | 20 | ja | 1
davon Eigenspektren | 22 |  |  | 4.4 | ja | 1
In Bezug auf Identifizierte: |  | Inklusivität
Positive Übereinstimmungen (PA) | 22 | Richtig-Positiv Rate | 100% | 95% | ja | 1
Negative Abweichungen (ND) | 0 | Falsch-Negativ Rate | 0% | 1% | ja | 1
Nicht-Ziel-Parameter
Zahl aller Spektren | 41
davon mit ID-Ergebnis | 39 | Identifiziert Rate | 95.12% | 30 | ja | 1
davon Eigenspektren | 39 |  |  | 7.8 | ja | 1
In Bezug auf Identifizierte: |  | Exklusivität
Negative Übereinstimmungen (NA) | 39 | Richtig-Negativ Rate | 100% | 99% | ja | 1
Positive Abweichungen (PD) | 0 | Falsch-Positiv Rate | 0% | 1% | ja | 1
Test auf spektrale Nähe | 41 | Anteil nah verwandter Sp | 0%
Von 22 identifizierten Proben des Parameters wurden unter Verwendung der vollständigen Datenbank 22 (=100%) richtig erkannt (Inklusivität). 0 (=0%) der identifizierten Proben des Parameters wurden falsch einer anderen Spezies zugeordnet.
In den 39 identifizierten Proben der Nicht-Ziel-Parameter wurden unter Verwendung der vollständigen Datenbank 100% richtig-negativ (Exklusivität) gezählt. Von diesen 39 Proben wurden 0 (=0%) fehlerhaft als Parameter identifiziert.
In den 41 Proben der Nicht-Ziel-Parameter wurden unter Verwendung einer nur den Parameter enthaltenden Datenbank bei 0 (=0%) Einträgen ein score > 2 erreicht.
Die Kriterien für die | Zielgerichtete Identifizierung | sind somit | erfüllt
Kommentar:
Freigabe an Standort | für (Abteilung / Bereich) | am (Datum) | von (Hdz.)
CVUA S | T, G, RK | 2023-02-14 | Rau
Das Verfahren ist für den vorgesehenen Einsatzzweck geeignet und zur Verwendung freigegeben.
als Prüfmethode für Parameter in LIMS eingepflegt
Eine unterschriebene Version dieses Validierungsreports befindet sich im erstellenden Labor.
Andere Keime, die in dieser Matrix vorkommen können.
Andere Keime, die auf diesem Agar wachsen können.
Vergleichsisolate/-materialien, die ähnliche Spektren liefern können (siehe Blatt "Ähnliche zu Parameter").
ja
nein